## Blood-Fluid Warmers
Supplier Reported Features: Blood/Fluid Warming Systems
Supplier Name | Product | Website Link | FDA | Description | Equipment | Indications for Use | Warming Method | Target Temperature | Maximum Flow Rates | Priming Volumes | Safety | Notes
3M (CN500236) | Ranger Blood and Fluid Warming SystemB4:F4 | https://www.3m.com/3M/en_US/company-us/all-3m-products/~/3M-Ranger-Blood-Fluid-Warming-Unit-Model-24500/?N=5002385+3294250939&rt=rud | K060939 (6/26/06) K082217 (10/6/08) | Consists of a warming device and an irrigation fluid warming disposable set. | Warming unit plus disposable single use warming set | Intended to deliver warm blood, blood products and fluids to adult and pediatric patients. | Dry Heat | 41°C | High Flow Set = 500 mL/min (30,000 mL/hr) | Standard Flow Set = 39‒44 mL High Flow Set = 150 mL Pediatric/Neonate Set = 20 mL | Audible and visible alarm for over-temperature (43°C). Audible and visible alarm for under-temperature (33°C). Independent back-up safety alert for 44°C to 46°C depending on unit that shuts off power to heating plates | Three flow sets: standard flow set, high flow set and pediatric/neonate set. Must be used in healthcare facilities by medical personnel trained in use of warming systems.
Smiths Medical (CN500237) | Hotline® Blood and Fluid Warmer | https://www.smiths-medical.com//en-us/products/temperature-management/blood-and-fluid-warming-systems/routine-flow-fluid-warming/hotline-blood-and-fluid-warmer | K911383 (10/17/91) K001764 (8/9/00) | Consists of a warming device and an irrigation fluid warming disposable set. Employs a recirculating solution heating system. | Warming unit plus individually packaged single use disposable warming set with a sterile fluid path. | Designed to warm blood and intravenous (I.V.) fluids and deliver them to the patient's intravenous access site at normothermic temperatures under gravity flow conditions. | Counter current fluid via triple lumen tubing. | 38°C | Up to 5,000 mL/hr | L-70 and L-70NI warming sets = 20 mL L-80 warming set = 21 mL | Visual and audible alarm for temperature 39.2°C or greater. Visual and audible "Check disposables" alarm indicates fluid warming set is not properly installed. Visual and audible "Add Recirculating Solution" alarm indicated reservoir level low. | Intended for use by trained medical personnel to provide routine flow of warmed I.V. fluid. Not to be used to warm platelets, cryo-precipitates or granulocyte suspensions. Multiple warnings for proper use within user manual. Requires reservoir disinfection and replacement of recirculating fluids. Maintenance checks with each use, every 30 days and every 12 months.
| Hotline3 Fluid Warmer | https://www.smiths-medical.com/en-us/products/temperature-management/blood-and-fluid-warming-systems/routine-flow-fluid-warming/hotline-3-blood-and-fluid-warmer | K911383 (10/17/91) | Consists of a warming device and an irrigation fluid warming disposable set. Employs a recirculating solution heating system. | Warming unit plus individually packaged single use disposable warming set with a sterile fluid path. | Designed to warm blood and intravenous (I.V.) fluids and deliver them to the patient's intravenous access site at normothermic temperatures under gravity flow conditions. | Counter current fluid | 41°C | 50 to 5,000 mL/hr | HOTLINE®3 warming set = 20 mL | Visual and audible "Over Temperature" alarm when recirculating solution is at or above 43.1°C. Visual and audible "Check disposables" alarm indicates fluid warming set is not properly installed. Visual and audible "Add Recirculating Solution" alarm indicated reservoir level low. | Intended for use by trained medical personnel to provide routine flow of warmed I.V. fluid. Not to be used to warm platelets, cryo-precipitates or granulocyte suspensions. Multiple warnings for proper use within user manual. Requires reservoir disinfection and replacement of recirculating fluids. Maintenance checks with each use, every 30 days and every 12 months.
| Level 1 H-1200 Fast Flow Fluid Warmer | https://www.smiths-medical.com/en-us/products/temperature-management/blood-and-fluid-warming-systems/fast-flow-fluid-warming-trauma/level-1-h1200-fast-flow-fluid-warmer-with-integrated-air-detectorclamp | BK020043 (6/16/03) | An intravenous fluid warmer with pressure chambers, air detection and automatic clamping capability where I.V. fluid and/or blood products are warmed by a sealed heated recirculating solution. | Warmer system and multiple D or DI series disposable administration sets. | Provides rapid flow warmed fluids (crystalloid or blood products) as volume replacement for patients suffering from blood loss due to trauma or surgery. | Counter current fluid | 41°C | DI-50 = 500 mL/min D/DI-60HL; D/DI-65HL = 530 mL/min D/DI-70; D/DI-75 = 750 mL/min D/DI-100; D/DI-150 = 950 mL/min D/DI-300; D/DI-350 = 1400 mL/min | DI-50 = 56 mL D/DI-60HL = 74 mL D/DI-65HL = 164 mL D/DI-70 = 70 mL D/DI-75 = 160 mL D/DI-100 = 65 mL D/DI-150 = 155 mL D/DI-300 = 90 mL D/DI-350 = 180 mL | Visual and audible "Over Temperature" alarm for temp at or above 43.9°C. Visual and audible "Air Detected" alarm when air is detected in the gas vent or filter assembly. Visual and audible "Disposables" alarm when assembly is improperly installed. Visual and audible "Add Recirculating Solution" alarm when reservoir is low. | Intended for use by trained medical personnel to provide routine flow of warmed I.V. fluid. Not for use in warming platelets, cryo-precipitate or granulocyte suspensions. Multiple warnings for proper use within user manual. D/DI-65HL, D/DI-75, D/DI-150 and D/DI350 disposable administration sets contain the F-50 gas vent/filer assembly which vents trapped air in the line to prevent it reaching the patient.
| NORMOFLO Irrigating System | https://www.smiths-medical.com/en-us/products/temperature-management/blood-and-fluid-warming-systems/irrigating-systems/normoflo-fluid-warming-device-irrigating-system | K873435 (10/5/87) K072080 (6/2/08) | Designed for safe, rapid inline warming of irrigation fluids as they are administered to patients | H-1100 warming irrigation system and disposable administration sets. | Provide in-line warming of irrigating fluids for the prevention of unintended periopertive hypothermia. | Counter current fluid | 41.7°C | IRI-600; IRI-600B, IR-700; PL-6 sets provide flow rates ranging from 620‒650 mL/min | Not available | Not available | Intended for use by trained medical personnel to provide in-line warming of irrigating fluids. May be used for a variety of procedures: laparoscopy, arthroscopy, hysteroscopy, wound debridement, urology, liposuction, lavage. Automatic IV pole lifts up to 12 L of fluid (four, 3L bags). Single disposable for one patient may be used for both the procedure and post-op irrigation, helping to reduce costs.
Belmont Medical Technologies | Belmont Buddy Lite/Belmont Buddy Lite AC | https://belmontmedtech.com/portable-iv-pump | K090508 (3/26/09) K112639 (10/25/11) | A portable, battery-powered, blood and fluid warmer used to prevent hypothermia during fluid administration in pre-hospital settings | A small heating system and single use disposable set. | Prevention of hypothermia during fluid administration in pre-hospital settings. | Dry Heat | 38°C ±2°C | 80 mL/min (input temp 20°C) 50 mL/min (input temp 10°C) | 7.5 mL | Not available | Warms fluid and/or blood close to the cannulation site. Automatic air removal to prevent air emboli. Warms up to 4.4 liters of fluid on a single battery charge. Belmont Buddy Lite AC extends warming capability when AC power is available.
| Belmont Buddy Rapid Infuser RI-2 | https://belmontmedtech.com/rapid-infusion-pump | K972284 (11/13/97) K983975 (11/30/98) K992672 (8/24/99) K032674 (9/16/03) K091855 (7/1/09) K141654 (7/18/14) | Provides high-speed delivery of warmed fluids. | Infusion pump and 3-spike disposable set. Additional optional disposables available based on needs. | High blood loss surgical procedures, trauma or where rapid replacement of warmed fluid is required. Infusion of warmed fluids to rewarm patients after surgery or those with hypothermia. Infusion of warmed fluid for irrigation in urology procedures. | Electromagnetic Inductive Heating | Flow ≤ 50 mL/min = 39°C ± 1°C Flow ≥ 60 mL/min = 37.5° ± 1°C | 1,000 mL/min | 100 mL | Monitors infusate temperature, line pressure, and air with alarms. Automatic air detection and removal. Automatic pressure monitoring and adjustment in flow for pressure > 300 mmHg. Low battery alarm. | Battery back up > 30 minutes (at 50mL/min without heat) to allow for patient transport. Not for use in warming platelets, cryo-precipitate or granulocyte suspensions. Not intended for drug administration. Do not add lactated ringers, dextrose in water and hypotonic sodium chloride solutions to blood components.
Biegler Medizin Electronik | BW685/BW685S | https://www.biegler.com/en/bw685 | K121198 (7/12/12) | A warmer for fluids and transfusions. A heat exchanger warms fluid as it continuously flows through an extension tubing. | Warmer system (± Tube Flow) and three options for extension sets | To warm infusions and transfusions to minimize risk of hypothermia in adult and pediatric patient populations. | Dry Heat | 38.5°C | Flow rates are neither monitored or directly controlled by the device. | Not available | Visual and audible high temperature alarm. Low temperature information signal. Three independent safety shutdowns. | Cannot be used as a flow rate monitoring or control system. Cannot be substituted for infusion pumps or as medication administration pumps.
EMIT Corporation | HypothermX™ LG | https://www.emitcorp.com/products.html# | Unable to find | Portable fluid warming device for intravenous fluid or blood products to prevent/treat trauma or hypothermia. Consists of a single unit placed between a standard IV infusion set and a standard IV extension set. | Fluid warmer unit | Therapeutic warming in the field with up to 1 hour continuous use. | Conduction via aluminum heat exchanger using a flameless hydrocarbon combustion process with isobutane/propane mixture | 38°C ± 3°C | 100 mL/min | Not available | Visual alarms for <35°C (blue), 42°‒46°C (yellow) and >46°C (red). Audible alarm for 42°-46°C (every 10‒20 seconds), >46°C (every 1‒2 seconds). Safety switch turns off at 50°C. | Warms up to 2 liters of saline (base temp 15°C) and 2 units of packed red blood cells (base temp 5°C) to 38°C ± 3°C with infusion rate between 50 mL/min to 100 mL/min and ambient temp 0°‒40°C.
| HypothermX™ HX100 | https://www.emitcorp.com/products.html# | K103801 (5/20/11) | Portable fluid warming device for intravenous fluid or blood products to prevent/treat trauma or hypothermia. Consists of a single unit placed between a standard IV infusion set and a standard IV extension set. | Fluid warmer unit | Therapeutic warming in the field with ≥ 4 hours continuous use. | Conduction via aluminum heat exchanger using a flameless hydrocarbon combustion process with isobutane/propane mixture. | 38°C ± 3°C | 200 mL/min | Not available | Visual alarms for <35°C (blue), 42°‒46°C (yellow), and >46°C (red). Audible alarm for 42°‒46°C (every 10‒20 seconds), >46°C (every 1‒2 seconds). Safety switch turns off at 50°C. | Warms up to 5 liters of saline (base temp 15°C) and 15 units of packed red blood cells (base temp 5°C) to 38°C ± 3°C with infusion rate between 50 mL/min to 200 mL/min and ambient temp 0°‒40°C.
Estill Medical Technologies | Thermal Angel™ | https://thermalangel.com/ | K984640 (7/1/99) | A portable, battery powered, disposable, single unit that is placed in-line between a standard IV drip set and an standard IV extension set. | Disposable fluid warmer unit and 9 inch IV extension set. Ultra and standard batteries, chargers and power cables. Optional Ultra or standard emergency packs for flight service. | To warm intravenous fluids and blood products prior to administration. | Dry Heat. A heating blanket covering stainless steel tubing. | 38°C | 200 mL/min | Not available | Visual alarms: No Power = red fading to dim/or no light: Decrease Heat = green fading to dim or no light. Safety switch turns off at 41°C. | For use by health care professionals in clinical settings and out in the field. Use with any standard IV or blood administration set. Placed near intravenous access site to reduce heat loss. Rapid warm-up time (within seconds)
Gentherm | ASTOFLO PLUS eco | https://www.gentherm.com/en/medical/normothermia | K082765 (12/18/08) | Universal fluid warmer that warms fluids without use of warming sets or proprietary disposables. | A control warming unit and a sleeve warmer. | To warm blood transfusions, intravenous fluids, dialysis lines, and irrigation fluids. Two tubing profiles: 4-5 mm and 6-7 mm | Dry Heat | Not found Heating Profile Temperature = 33°C to 43°C | Various | Not available | Safety temperature cut off: 43.6°C ±0.5˚C Low temperature alarm: 3°C below set temperature for longer than 10 minutes. | Uses standard intravenous/blood administration sets. Warms fluid up to patient cannulation site. Waterless, dry-heat warmer that can be wiped clean and disinfected in seconds. Only appropriately trained and qualified healthcare providers and staff should use this product.
| ASTOTHERM PLUS | https://www.gentherm.com/en/medical/normothermia | K991160 (9/2/99) | Electronically powered, dry warmer that applies heat to intravenous administration sets. | A control warming unit, a sleeve warmer and a sterile extension line (disposable). | Prevention of hypothermia, warming of transfusions, infusions, dialysis, hemofiltration, and apheresis. | Dry Heat | Three temperature selections: 39°C; 41°C; 43°C | Various | Two available extension lines: 4 mm diameter (320 cm L) = 22 mL 4 mm diameter (575 cm L) = 40 mL 6.8 mm diameter (490 cm L) = 89 mL | Three independent safety alarms for over or under heating. A circuit shut-off to prevent over-heating. | Warm up time less than 1 minute. Warms fluid up to patient cannulation site. Waterless, dry-heat warmer that can be wiped clean and disinfected in seconds.
Smission-Cartledge Biomedical | ThermaCor® 1200 | https://www.thermacor1200.com/ | K052055 (10/26/06) | Portable (tabletop/pole-mount) warming device for the use of warming intravenous fluids/blood or irrigations. | Heating system/infusion pump with sterile disposable kit. | Slow to rapid infusion of warmed fluids or blood products to treat blood loss from trauma/surgery; to prevent hypothermia; to rewarm after surgery; or to irrigate during urological procedures. Adult and pediatric patients. | Dry Heat | 40°C | 1200 mL/min | Not available | Monitoring of temperature, air and pressure. Automatically reduces flow based on selected pressure limits (100 or 300 mmHg) | Battery back up (30 minutes at maximum flow 1200mL/min; up to 2 hours) without heat to allow for patient transport. Has bolus mode capability.
Vyaire Medical | enFLow Fluid Warming System (Disposable cartridge under global recall: model numbers 980200EU, 980202EU) | https://www.vyaire.com/news-events/vyaire-medical-announces-global-market-recall-enflowtm-fluid-warming-system-disposable | K060537 (6/8/09) K112902 (3/14/12) K121775 (11/28/12) K130867 (6/14/13) | Consists of a warmer, controller (power supply), and singe use sterile disposable cartridge to deliver warmed infusate up to 40°C at flow rates of 1mL/min to 200 mL/min. | Warmer, controller and single use disposable cartridge. Optional intravenous tubing extension set. | For warming blood, blood products and intravenous solutions prior to administration to prevent hypothermia. | Dry Heat | 40°C | 200 mL/min | Not available | March, 2019: Global marketing recall of the enFlow™ cartridge due to potential to elute aluminum during intravenous warming therapy with fluid and blood administrations. Customers should suspend all use of the enFLow™ system at this time. | For use by healthcare professionals in the hospital, in clinical settings and in the field. Not to be used with platelets or drugs.
Note: This document is for informational purposes only and is not intended to replace individual clinical decision-making, which is the sole and independent responsibility of the practitioner. HealthTrust expressly disclaims any liability for treatment decisions. Please direct any questions or comments to clinical.research@healthtrustpg.com, or to your supplier representative. ©2021 HealthTrust. All Rights Reserved.
[7 empty] | Initial or Update | Date | Completed by | Changes Made
Updated
[7 empty] | Updated | 02.2021 | FP | Formatting
[7 empty] | Initial | 01.2021 | HAM | Created