## Active Body Surface Warming
Supplier Reported Features: Patient Temperature Management
Supplier Name | Product | Website Link | FDA Approval | Method of Heating | Description of Device | Indications for Use | Warming System Components | Modes or Settings | Temperature Range | Patient Temperature Monitoring | Safety | Notes
Pintler (CN 7154) | Pintler Patient Warming System | http://www.pintlermedical.com/ | K113705 | Conductive | Thermal regulating pad with a power control unit mounted on a standard IV pole. Utilizes low voltage resistant heating element, DC power. PEAK: Warming OR table set; SUMMIT: Orthopedic or Bean Bag use; VISTA: Heated stretcher pad cut to fit all stretchers in perioperative, emergency, diagnostic, endoscopic, and other clinical environments | To maintain normothermia during perioperative experience for pediatric and adult patients | 3.5 inch mattress, ribbon cable, and power control unit. Single-use thermal reflective booties and mittens also available | Standby and Warming ‒ with temperature selections between 37°C to 42°C | 37°C to 42°C | No | Automatic shut off if temperature goes above 42°C | ● Heated pads are also pressure reduction pads ● Foam in pad is anti-microbial and anti-bacterial ● Pads can be disinfected and reused ● Able to X-ray through the pad ● Able to be used on any type of operating room table ● SCIP compliant ● Pads last 2-3 years, repair is kit included, external stretch vinyl cover protection
Augustine Surgical Inc. (CN 2600) | HotDog® Patient Warming System | https://www.augustinesurgical.com/ | K201779 K112488 K094056 | Conductive | WC77 Temperature Management Controller- Multiport: Provides power to heat warming devices to a controlled temperature. Ability to use up to 4 warming blankets and one mattress. WC71 Temperature Management Controller - Single Port: Provides power to heat warming devices to a controlled temperature. Single port for one warming device. Primary use is preoperative/postoperative setting | To prevent or treat hypothermia and to provide warmth to adult or pediatric patients. For Use in hospitals and surgical centers including operating, recovery, emergency rooms and med/surg units | Temperature control unit, reusable warming devices (warming blanket or pressure-relief mattress), and connecting cable | ● Manual warming mode ● Auto Mode: Automatic control and adjustment of the warming devices based on the patient's temperature. For use when control unit is monitoring patient temperature | Blanket ports: 37°C to 43°C ± 1.0°C Mattress ports: 35°C to 40°C ± 1.0°C Adjustable in 1°C increments | Yes | ● Automatic shut off feature preset for six hours of non-interaction. Time interval can be adjusted ● Alarms for over-temperature (primary and secondary), failure to reach temperature within 10 minutes, port current reached, sensor or cable failure, device fold detection, auto-mode disengaged, and hardware failure ● Warming devices will stop warming if an alarm appears | ● Attached to IV pole ● Mattresses and blankets can be disinfected and reused. Several disinfectants are compatible ● Do not warm ischemic or non-perfused tissue as it may result in thermal injury ● Do not warm patients receiving transdermal medication as it increases medication delivery ● Do not use the warming mattress with other under-patient thermal management systems
VitaHeat Medical LLC (CN 20299) | VITAHEAT™ Patient Warming System (Note: Filed Chapter 7 Bankruptcy on 12-21-18 and Noted Out of Business February 1, 2019) | No Website | K132454 | Conductive | The VITAHEAT™ Patient Warming System is clinician or patient-controlled, and consists of a controller, a reusable warming mattress, single use disposable covers, and accessories. It is powered by a rechargeable 18V lithium ion battery. A non-heating thigh pad is included that is used for thigh support | To prevent or treat hypothermia and to provide warmth to patients in situations where the patient may not maintain a state of normothermia | Controller, reusable warming mattress, single use disposable covers, non-heating thigh pad, rechargeable 18V lithium ion battery, charging unit, detachable extension cord | ● Manual mode controlled by the patient or clinician ● Temperatures are adjusted depending on the power setting selected by the user (higher and lower buttons, when pressed, raise or lower temperature delivery) | 35°C to 38.5°C ± 1.5°C | No | Automatic system shut off for warming mattress temperature > 41°C | ● Powered by a rechargeable 18V lithium ion battery (separate charging unit and second rechargeable ion battery provided with device) ● Portable: Battery life = 3 to 5 hours maximum when fully charged ● Controller is placed on hospital bed rail, stretcher rail, or on the patient bed
Encompass (CN 5251) | Thermoflect® Passive Warming Products | https://www.encompassgroup.com/patient-warming | Not applicable | Reflective | Thermoflect® is a four-layer constructed design to promote normothermia by using heat reflective technology and quilted channels to capture and reflect radiant body heat while preventing convective heat loss | For use with adult and pediatric patients to keep patients warm and comfortable in the prehospital, hospital and transfer settings | Multiple components available: Thermoflect Transport Cocoon, Patient Transfer Kit, Staff Apparel, Total Access Gown, IV Access Gown, Blankets (S, M, L), Adjustable Leggings, Bouffant Cap, and Patient Warming Admission Kits | Not applicable | Not applicable | No | Non-conductive | ● Start working immediately after application ● Do not require any equipment, electrical or other consumables ● Total Access Gown has a 360° quick-release strip at the torso to transform the gown into two pieces ● Thermoflect Transport Cocoon has full-body super absorbent inner core to contain fluids
3M (CN 735) | Bair Hugger™ Warming Blanket System Bair Hugger™ Warming Gown System Bair Hugger™ Temperature Management System Bair Hugger™ Temperature Monitoring System | https://www.bairhugger.com/3M/en_US/bair-hugger-us/ | K171373 K041686 K021473 | Convective | The temperature management unit maintains normothermia by delivering warmed, room temperature, or cooled air directly to a Bair Hugger blanket or gown via a flexible hose. Small perforations in the patient-side of the blanket disperse the air over the patient | Warming Systems are used to prevent and treat hypothermia in adult and pediatric patients. The Model 775 Temperature Management System is also used to promote normothermia in situations where the patient is too warm | Temperature management unit, multiple over-the-body blankets, specialty and underbody blankets, pediatric blankets, warming gowns, and accessories (replacement power cord, hose, filter, sheet clip, rolling cart and handle, and bedrail mount) | ● Bair Hugger Model 775 temperature management unit has standby and 4 operating modes: Ambient, 32°C, 38°C, 43°C ● Two fan speeds: High (default) and Low | 32°C to 43°C Ambient setting Low and High Fan settings | Yes, via the temperature monitoring system: a noninvasive continuous measurement of core temperature | ● Over temperature visual and audible alarm at 53°C ± 3°C with automatic cessation of blower and heater ● Several nonhazardous fault visual and audible alarms, resulting in cessation of blower and heater | ● Do not apply heat to lower extremities during aortic cross-clamping as injury may occur when heat is applied to ischemic limbs ● Filtration system is a high efficiency 0.2 µm air filter. Replace air filter every 12 months or after 500 hours of use ● Noise level 51 dBA
SourceMark (CN 76177) | KoZee™ Convective Warming System | https://www.sourcemarkusa.com/product/kozee | K162679 K190221 (Under IOB Medical) | Convective | A low volume air-flow system comprised of a proprietary network of digital sensors coupled with a ceramic heating module that operates at 28 cfm, approximately 41% less airflow requirement to maintain normothermia | Indicated for hypothermic patients or normothermic patients for whom induced hyperthermia or localized increase in temperature is clinically indicated | Heating system, connecting hose and single use sterile blanket. Optional hose sleeve for easy cleaning | ● Ambient, 32°C, 38°C, 43°C ● Two air flow speeds: High and Low flow | 32°C to 43°C Ambient setting Low and High Fan settings | No | ● Automatic shut off at 47°C ± 2°C | ● Dual layer filtration with HEPA specifications of 0.2µm filtration/99.999% efficiency ● Low volume air flow and blanket with greater permeability ● Quiet device ● 25 blanket styles
Stryker (NC) (The 37° Company) | Mistral-Air | https://www.stryker.com/us/en/portfolios/medical-surgical-equipment/temperature-management.html | K171234 K101705 K073712 | Convective | The Mistral-Air® Warming System is a forced air warming device comprised of a warming unit and a variety of blankets intended to raise and maintain patient temperature by means of surface warming. Only to be used with disposable single use Mistral-Air® blankets | A forced air warming device used to raise and maintain patient temperature by surface warming | Control unit, flexible hose, warming blanket (sterile or non-sterile) draped over the patient. Optional basket cart | Manual selection of four settings: Ambient (Fan only) Low (32°C) Medium (38°C) High (43°C) | Ambient 32°C 38°C 43°C | No | ● Visual and audible primary (>45.5°C) and secondary (>56.4°C) over temperature safety alarms ● Visual and audible under temperature safety alarm set at 6°C under the set temperature ● Visual and audible alarms when filter change is required ● Audible alarm suppression ● Visual and audible service light indicator alerting user for ≥2,000 hours | ● Do not apply heat directly to open wounds ● Do not apply warming system to ischemic limbs ● Noise level 54‒55 dBA
Inspiration Healthcare (NC) | AlphaCore Patient Warming System | https://www.inspiration-healthcare.com/products/patient-warming/alphacore | Not found in FDA 510K database | Conductive | A patient warming system that provides controlled warming to maintain normothermia for neonatal, pediatric and perioperative environments | To maintain normothermia, or treat hypothermia in the adult, pediatric or neonatal patients | AlphaCore controller, mattress, connection cables | User selectable range of temperature settings: 25°C to 40°C | 25°C to 40°C | No | Not found | ● Universal mattress fits 95% of operating room tables ● Able to connect several warming devices at the same time ● Silent operation ‒ reduces background noise
Halyard Health (NC) | Halyard Health Patient Warming System | https://products.halyardhealth.com/surgical-solutions/temperature-management.html | Not found in FDA 510K database. | Conductive | Unit draws temperature controlled water through disposable gel pads applied directly to the patient's body | Any condition where patient temperature control within a range from mild hypothermia to normothermia is required | M1000 System Control Unit, kits and components, gel pads | ● Automatic mode that analyzes the patient's rate of temperature change ● Customizable mode | 33°C to 37°C | Yes | ● Computer controlled algorithm avoids warming patients above normal range and prevents temperature spikes ● Alarm information not provided | ● Thin, comfortable gel pads allow placement on nearly any body surface ● Pads are non-sterile ‒ do not place in operative field. Carefully remove from patient's skin at completion of use. Do not allow water to contaminate surgical field when disconnecting ● Precise temperature management ± 0.5°C ● Do Not use on patients with ulcerations, burns, hives, rash, or history of skin allergy/sensitivity
Stihler Electronic Gmbh (NC) | Astopad® Patient Warming System | https://www.gentherm.com/en/medical/normothermia/astopad | K202197 | Conductive | The ASTOPAD Patient Warming System is a thermal regulating system which includes a connection cable that attaches to reusable conductive warming blankets for patient warming in clinical environments | Intended to prevent or treat hypothermia and provide warmth to adult and pediatric patients. Indicated for use in all areas of healthcare facilities where maintenance of normothermia is needed | Control unit (DUO310-NA) and one or optionally two added blankets (warming blankets/COVXXX-NA) with different sizes to fit different applications | ● Standby mode ● On mode with customizable temperature targets | 32.0°C to 39.0°C Able to adjust temperature in 0.5°C increments | No | ● Audible and visual alarm for overheating (actual temperature ≥1°C from set temperature) ● Audible and visual low temperature alarm (actual temperature is ≤1.0°C from set temperature) ● Additional overheating alarms for actual temperature ≥41°C and applied part temperature ≥45.0°C ● Audible and visual alarm for heater defect and sensor defect, with power shut off | ● Warming blankets can be used over or under the patient ● Use of 8 temperature sensors for blanket temperature control and monitoring ● Unit can be fitted with a battery for portable operation up to 2 hours ● System is "defibrillation proof" ● Do not use distal to arterial occlusion or on ischemic extremities ● Do not use on patients receiving transdermal medication as it increases medication delivery
Medical Solutions, Inc. (NC) | Cocoon Convective Patient Warming System | https://www.msi-healthcare.com/convective-patient-warming-system/ | K181671 | Convective | A microprocessor-controlled device that delivers a continuous flow of temperature-controlled air through a flexible hose to a warming blanket | Indicated for hyper- or hypothermic patients or normothermic patients for whom induced hyper- or hypothermia or localized temperature therapy is clinically indicated. May be used with adult or pediatric patients | The Cocoon Convective Warming System includes the Convective Warming Machine (CWS4000) and the Cocoon Disposable Patient Warming Blankets | Six settings: Ambient 34°C 37°C 40°C 43°C 46°C boost option | Ambient to 46°C | No | ● If setting 46°C chosen, system automatically reduces temperature to 43°C after 10 minutes ● Dual thermostats with high temperature cut off that stops therapy if unit overheats | ● 20 different blanket products ● Increased surface area of filter to extend time between filter changes ● Includes hose slip cover that facilitates ease of cleaning ● Sound level not provided; manufacturer claims "quietest unit in the market" ● Durable but soft blankets with universal tie system to secure blanket to unit
Smiths Medical (NC) | Level 1® Convective Warming System | https://www.smiths-medical.com/level-1 | K181699 | Convective | Consists of a high-flow convective warmer with hose-end temperature control that provides 4 temperature settings, including an ambient option. The warmer is used with a convective warming blanket and accessories | Intended to prevent or treat hypothermia when temperature therapy is clinically indicated, and to provide thermal comfort for patients. Can be used on adult or pediatric patients | Level 1 Convective unit, warming blankets and accessories | ● Standby Mode ● Four customizable temperature settings that include ambient, 37°C, 40°C and 43°C | Ambient to 46°C | No | ● Visual and audible alarms for overheat, maintenance needed and occlusion ● Visual and audible alarms for over-temperature: +1°C over selected temperature and/or above 45°C ● Visual and audible alarm for under-temperature: -1°C under selected temperature and/or below 36°C ● Safety circuit provides an independent means of shut off that discontinues power to the heater | ● Precise temperature control within ± 1°C ● Use of HEPA air filter ● Perforations on patient side of blanket gently disperses warm air over the patient ● Sound level is < 42 dBA
NovaMed USA (NC) | Koala Warming System | http://www.novamed-usa.com/koala_warming_home.html | Not found in FDA 510K database | Conductive | Reusable underbody warming with a pressure relief multi-layer mattress that straps directly to the OR table to provide transfer of thermal warmth to the patient | Any medical situation where there is a potential risk of hypothermia and unhindered patient access is a requirement. Recommended for open-heart, organ transplant, major abdominal, trauma, and orthopedic procedures. Applicable in OR, Critical Care Units, PACU, and ER. Suitable for adult, pediatric, neonatal patients | Controller and warming mattress that is available in 4 lengths. Overbody blankets also available | Not reported | 37°C to 40°C | No (NovaMed does have a separate temperature monitoring system that is not part of the Koala Warming System) | Not found | ● X-ray and radio translucent ● SCIP-INF-10 compliant ● Mattress lasts for 5+ years ● Noise Free
Gentherm Medical (NC) (Cincinnati Sub-Zero Products, LLC) | WarmAir® Convective Warming System | https://www.gentherm.com/en/medical/normothermia/warmair | K123946 | Convective | A small and compact warming unit designed to supply air at temperatures of ambient, 32.2°C, 37.8°C or 43.3°C to a patient via an air-distribution device | To prevent hypothermia and/or reduce cold discomfort before, during and after surgical procedures through convective heat transfer from the controller to a warm-air-heated blanket. For use with adult, pediatric or neonate patients | Temperature control unit, FilteredFlo® single-use disposable blanket, warming tube | Fan only, Low, Medium, High modes that correlate with the respective temperature settings | Ambient 32.2°C 37.8°C 43.3°C | No | ● Maximum temperature 43.3°C + 4°C ● Primary over-temperature limit at 52.0°C ± 3°C with audible/visual alarms and heater/blower shutdown ● Secondary over-temperature limit at 64.0°C or less with power shutdown ● Under-temperature limit at 29.4°C or less with audible/visual alarms and heater/blower shutdown | ● Close monitoring of patients is required when treating the following patients with high temperature setting: significant peripheral vascular disease; occlusive or diabetic in nature; low cardiac output; marginal cutaneous perfusion ● Do not apply heat to lower extremities during arterial cross-clamping. Thermal injury may occur if heat is applied to ischemic limbs ● Transparent to X-ray and imaging systems ● Information on noise level (dBA) not provided
| Norm-O-Temp Hyperthermia System | https://www.gentherm.com/en/medical/normothermia/norm-o-temp | K120081 | Conductive | A water re-circulating system providing either hyperthermia or normothermia treatment, or to keep a patient comfortable by maintaining blanket/pad water temperature through a conductive heat transfer. Water is heated and pumped from the device, through connecting tubing, to disposable or reusable blanket/pads, which can rest under, above or around the patient | Intended to prevent hypothermia during surgical procedures and promote comfort in the perioperative period. For use with adult, pediatric and infant (neonate) patients in operating rooms, post-anesthesia care units, intensive care units, and emergency rooms | A heater, circulating pump and blanket/pads | Manual water temperature settings for the circulating water in increments of 1°C | Water temperature range: 32°C to 42°C | No | ● Primary high temperature visual/audible safety alarm if water reaches 43.5°C ± 0.6°C ‒ pump and heater are turned off ● Secondary high temperature safety visual/audible alarm if water temperature reaches 44.5°C ± 0.6°C ‒ pump and heater turned off ● Independent high temperature safety alarm at water temp of 44.5°C ± 6°C with visual/audible and shut off of pump and heater ● 1°C overset alarm when water temperature exceeds set point temperature by 1°C ± 0.6°C. Audible/visual alarm and heater turns off; pump continues ● Audible/visual alarms for low water, defective water temperature sensor, power failure | ● Two blanket connections. Use Gentherm recommended blankets and pads ● Thermal injury can occur if heating/cooling therapy is applied to ischemic limbs ● Do not use warming system distal to arterial cross clamping ● Warming of transdermal medication patches can increase drug delivery and possibly injure the patient ● Use sterile water or water that has been passed through a filter less than or equal to 0.22 microns only. Do Not use De-Ionized water or distilled water ● Water leaks present a risk of infection ● Lifetime of unit is 10 years
Adroit Medical (NC) | ThermalX Heat Reflective Blankets | https://adroitmedical.com/reflective-warming/ | Not applicable | Reflective | Blankets and apparel made of heat reflective material that naturally preserves body temperature and prevents convective heat loss | For the prevention of patient heat loss in the operating room, PACU, and during transport. Suitable for all types of surgeries and emergency rescue | Blankets, caps, jackets, leggings, mittens, booties, papoose and kits | Not applicable | Not applicable | No | ● Non-conductive ● Do not use in MRI | ● Latex free, radio-translucent ● Safe for X-ray equipment and with laser surgery
| Soft-Air OR-110 | https://adroitmedical.com/forced-air-patient-warming-2/ | K073712 | Convective | A convective air warming system that consists of a warm air blower and soft air blankets | Provides safe and controlled surface warming to assist adult patients maintain normothermia. Designed for the operating room, recovery room, anesthetic room, ICU, Med-Surg floors, in-hospital patient transport and emergency rooms | Warming unit, hose, blankets and accessories | Four settings based on selected temperature | Ambient 32°C 38°C 43°C | No | Not found | Double filtration using a premium HEPA filter at the blower unit and proprietary nonwoven blanket material that filters air a second time at the patient
Steris Healthcare (NC) | STERIS Patient Warming System | https://www.steris.com/healthcare/products/surgical-table-accessories/patient-warming/patient-warming-system | K200446 | Conductive | A modular conductive patient warming system to be used during the pre-, intra- and postoperative stages of surgical procedures. For use in operating rooms, recovery rooms and emergency rooms | To provide heat for treatment of hypothermia, maintenance of normothermia, and thermal comfort for adults and pediatric patients | Temperature control unit, pressure-relieving warming mattress and optional over-body blanket | Five settings from 35°C to 39°C | 35°C to 39°C | No | Not found | Not intended for use in infant/neonatal populations
Medtronic (NC) | WarmTouch 6000 | https://www.medtronic.com/covidien/en-us/products/temperature-management/warmtouch-wt6000-warming-unit.html | K123083 | Convective | An electro-mechanical blower that delivers heated air from the blower's flexible hose to a lightweight blanket that is draped over a patient. Warmed air reaches the patient through small perforations in the blanket | Intended for the prevention and treatment of hypothermia, and for the management of normothermia in the preoperative to postoperative setting. Blankets available for adult and pediatric patients | Warming unit, blankets and accessories | Five settings, including ambient and boost | 5 temperature settings: Ambient 34°C 40°C 45°C 47°C | No | ● Onscreen filter expiration warning ● No information on alarms | ● Latex free ● Average time for contact surface temperature to rise from 23.0°C to 37.0°C = 6 minutes ● No information on noise level (dBA)
Note: This document is for informational purposes only and is not intended to replace individual clinical decision-making, which is the sole and independent responsibility of the practitioner. HealthTrust expressly disclaims any liability for treatment decisions. Please direct any questions or comments to clinical.research@healthtrustpg.com, or to your supplier representative. ©2021 HealthTrust. All Rights Reserved.
|  |  |  | Initial or Update | Date | Completed by | Changes Made
|  |  |  | Updated | 03.2021; 04.2021 | FP | Formatting
|  |  |  | Initial | 02.2021 | HAM | Created

## Targeted Temperature Management
Supplier Reported Features: Patient Temperature Management
Supplier Name | Product | Website Link | FDA | Method of Cooling/Heating | Description of Device | Indications for Use | Warming System Components | Modes | Device Temperature Range | Patient Targeted Temperature Range | Patient Temperature Monitoring | Safety | Notes
Stryker (CN 500208) | Altrix Precision Temperature Management System | https://www.stryker.com/us/en/acute-care/products/altrix.html | K152266 K171149 K180834 | Conductive (Body Surface) | Supplies water to an individual device or multiple thermal transfer devices simultaneously with each circuit monitored separately. Unit uses a patient temperature probe to provide closed loop feedback for automatic patient temperature management and monitoring | To maintain a pre-set body temperature as determined by the physician; maintain normal body temperature during surgical procedures; temperature reduction in patients where clinically indicated; for use with adult and pediatric patients | Controller, reusable hose sets, thermal transfer devices (blankets, vests, leg wraps), patient temperature probes, reusable adapter cables, reusable patient temperature output cable | ● Automatic mode cools or warms the patient to a selected target temperature. Three cooling rates are available and four warming rates are available in the automatic mode. Patient temperature controlled within ± 0.3°C of target temperature ● Manual mode where only the water temperature is controlled. A temperature probe is not required. Once water target temperature is achieved, the water temperature is controlled ± 3°C ● Monitor mode: No therapy is delivered, only patient temperature is monitored | 4.0°C to 40.0°C | 32°C to 38°C Patient temperature controlled within ± 0.3°C of target. Cooling and warming rates options (Max, Med, Minimum, Custom) | Yes Adhesive skin, 9FR and 12FR general purpose probe, 14FR and 16FR foley catheter temperature sensing probe options | ● Contraindicated for: Primary or secondary Raynaud's Phenomenon; application to lower extremities distal to aortic cross-clamping ● Multiple safety alarms, including but not limited to: remove from use, power loss, check patient probe, patient temperature deviation, water temperature deviation, check water flow, no water flow, no water, therapy pause time out, normothermia deviation and low battery | ● Use sterile distilled water or water that has been passed through a filter ≤ 0.22 microns ● Expected service life of 5 years under normal use with appropriate maintenance ● Disinfect external surfaces before use. Disinfect the internal water circuit and hoses before first use, every 14 days, and before storage with BruClean TbC disinfectant tablets
Cincinnati Sub-Zero/Gentherm (CN 500281) | Blanketrol® III Hyper-Hypothermia Temperature Management System | https://www.gentherm.com/en/medical/hyper-hypothermia/blanketrol-3 | K110104 K101589 | Conductive (Body Surface) | Temperature management system that heats or cools circulating water through disposable or reusable pads/blankets to maintain a desired patient temperature through conductive heat transfer | To lower or to raise a patient's temperature and/or maintain a desired temperature through conductive heat transfer. Blanket rests under and/or on top of the patient. Can be used with adult, pediatric and neonatal patients | Temperature control unit, Maxi-Therm Lite Blanket, PlastiPad Blankets, Kool-Kit (including head wrap), connecting hose (unless PlastiPad with hose attached), 400 Series Probe/connecting cable, distilled water | ● Manual mode for pre-conditioning temperature prior to therapy start, fever control or normothermia without continuous temperature monitoring ● Auto Control Mode for shivering, larger patients with dense body mass, patients with higher metabolic rates, perioperative normothermia ● Gradient 10C SMART Mode for neonates and small patients, those who overshoot target temperature, very sedated patients. Maintains water temp at maximum of 10°C from patient's temperature reading ● Gradient Variable Mode with Smart for neonates and small patients, those who overshoot target temperature, very sedated patients. Allows provider to customize temperature offset ● Monitor Only Mode shows the patient's temperature reading, but does not cool/warm or circulate the water. Used to monitor post-therapy | 4°C to 42°C | 30°C to 40°C | Esophageal, skin or rectal probes for temperature monitoring (400 series temperature probes) | ● If water temperature reaches 44°C ± 2.0°C, automatic shut off occurs and visual/audible High Limit alarms indicating to remove from service. Back-up alarm alerts in the event of a primary alarm failure ● If water temperature reaches 2.0°C ± 2.0°C, automatic shut off occurs with visible/audible Low Limit alarms indicating to remove from service. Back-up alarm alerts in the event of a primary alarm failure ● Additional alarms: check probe, bad probe, low water, low flow, check flow switch, preventative maintenance required | ● Uses distilled/sterile distilled water ● Exceeding 40°C for extended periods can cause tissue damage and burns ● Do not use distal to cross clamping as thermal injury may result ● Do not apply to ischemic limbs as thermal injury may occur ● Fluid circuit disinfection/dry storage procedure should be performed quarterly at a minimum
BD/Medivance (CN5288) | Artic Sun | http://www.medivance.com/ArcticSun5000 | K002577 K071341 K101092 K142702 K161602 K200225 | Conductive (Body Surface) | A non-invasive thermal regulating system indicated for monitoring and controlling patient temperature within a range of 32°C to 38.5°C in adult and pediatric patients | Indicated for monitoring and controlling patient temperature in adult and pediatric patients of all ages | Arctic Sun 5000 Control Module, disposable non-sterile ArcticGel Pads, and a commercially available medical temperature probe (naso-pharyngeal, bladder, rectal or esophageal) | ● Normothermia mode if goal is to maintain a patient temperature at a pre-defined target temperature for an indefinite period of time ● Hypothermia mode to reduce and maintain a patient temperature at a set target temperature for a defined period of time, then slowly re-warm the patient at a controlled re-warming rate ● Options for New Patient or Continue Current Patient | 4°C to 42°C | 32.0°C to 38.5°C Patient temperature controlled within 0.4°C of target | Yes Any commercially-available Yellow Springs Instrument 400 Series (YSI 400) compatible patient temperature probe can be used | ● Alarms: Audio/visual signal that repeats every 10 seconds until the alarm is cleared. Notifies user that a condition may potentially be unsafe to the patient or the device. Main alarms include high or low patient temperature, high or low water temperature, system self-test failure ● Alerts: Audio/visual signal that repeats every 25 seconds ● For both alarms and alerts, the screen will display the description of the problem that triggered the alert and solutions/instructions on troubleshooting | ● Use sterile water ● Use only with ArcticGel pads (non-sterile and single use) ● Flow 5 liters/minute ● Cooling 37°C to 33°C: approximately 2 hours ● Warming (max) 33°C to 37°C: approximately 6 hours
BrainCool (CN 59003) | IQool™ System | https://braincoolinc.com/ | K162523 K180375 | Conductive (Body Surface) | A non-invasive thermal regulating system indicated for monitoring and controlling patient temperature within a range of 32°C to 38.5°C. | A temperature regulating system indicated for monitoring and controlling patient temperature | Refrigeration and control unit, five liters of diluted monopropylene glycol cooling fluid, cooling pads (2 thigh, 2 torso, 2 optional for head and neck), stabilization insulation, and accessories (USB flash drive, filling pitcher) | ● Normothermia: Controls patient temperature ● Hypothermia: Cool patient and Rewarm patient | 4°C to 40°C | 33°C to 38.5°C Patient temperature controlled within ± 0.4°C | Yes Via one or two commercially available third-party YSI 400 series temperature probes (esophageal, foley or rectal) | ● High and Low coolant temperature limit ● Additional alarms not noted | ● Pads prefilled, no purging ● Coolant flow rate = 4.5 liters/min. ● Cooling rate of 1.38° to 1.61°C/hr.
Zoll (CN 2599) | Thermogard XP® Temperature Management System | https://www.zoll.com/products/temperature-management/thermogard-xp-tms | K150046 K153226 | Conductive (Intravascular) | An integrated electro-mechanical heater/cooler that consists of a temperature monitor, a temperature controller unit, a heat exchanger unit, and roller pump. It supplies temperature controlled sterile saline to the indwelling catheter that is placed percutaneously in the patient | ● ICY™, Quattro,® and Solex 7™ catheters are used to achieve and or maintain normothermia during cardiac surgery/recovery and induce/maintain/reverse mild hypothermia in neurosurgery patients in surgery/recovery (ICU) ● Cool Line™ and Solex7™ catheter for use in fever reduction, as an adjunct to other antipyretic therapy, in patients with cerebral infarction and intracerebral hemorrhage who are intubated/sedated and require CVL | External heat exchange console and an endovascular heat-exchange catheter connected via a sterile heat exchanger/tubing circuit (Start-Up kit). Catheter and start-up kit are single-use disposables | ● Standby mode ● Run - Treatment Modes: Max Power (console aims to make the patient temperature the same as the selected target temperature); Controlled Rate (console attempts to move the patient's temperature to the target temperature at the programmed rate of heat exchange °C/hr. Available in both cooling and warming settings); Fever (console will start cooling the patient once the patient's temperature is above the target temperature) ● Ability to set cooling or warming rate between 0.10°C/hr and 0.65°C/hr | 0°C to 42°C | 31°C to 38°C Patient temperature accuracy within ± 0.2°C | Yes Via an indwelling YSI 400 thermistor temperature sensor | ● Visual and audible alarms for High and Low patient temperature, air trap fault, coolant low or empty, primary or secondary temperature probe dislodged/disconnected, pump failure or door open | ● Lifetime of catheters vary by catheter and placement location, typically 4 to 7 days. Upper-body catheters include Solex 7™ and Cool Line.® Lower-body catheters include Quattro®, Icy®, and Cool Line® ● Catheters are 9.3 F ● Requires sterile saline ● Pump capacity = 7 liters/min at pump head
Attune Medical (CN 47665) | ensoETM | https://www.ensoetm.com/ | K172493 K172029 K180244 K180742 | Conductive (Esophageal) | A single-use, multi-lumen silicone tube that is inserted into the esophagus to control patient temperature. It is connected to an external heat exchange unit. Positioned within the esophagus near the great vessels and left atrium of the heart, water circulates inside a closed loop system to cool or warm a patient | A thermal regulating device intended to connect to a temperature management control system to control patient temperature, allow administration of enteral fluids and provide gastric decompression. Compatible temperature control systems include: Gaymar Medi-Therm III Conductive Hyper/Hypothermia System, Stryker Altrix Precision Temperature Management System, Cincinnati Sub-Zero Blanketrol II or III Hyper-Hypothermia System or Norm-O-Temp Hyperthermia System | A single-use, silicone tube, heat exchange unit (compatible with Altrix,® Medi-Therm® III, Blanketrol® II or III, Norm-O-Temp®) | Based on temperature control unit used with the device | Based on temperature control unit used with device | Based on temperature control unit used with device | Based on temperature control unit used with device | Based on temperature control unit used with device | ● Method bypasses the skin's shivering sensors, allowing more precise cooling ● Stomach access is maintained to allow for suctioning, feeding and medicine administration ● Intended duration of 72 hours
Inspiration Healthcare (NC) | Tecotherm NEO | https://www.inspiration-healthcare.com/products/neonatal-intensive-care/thermoregulation/tecotherm-neo-us | K140078 | Conductive (Body Surface) | Uses circulating fluid to provide controlled cold and heat treatment procedures and application of specific cold and heat doses to neonate and babies. Cold and heat are applied via mattresses/wrap to total body, body parts and areas depending on therapy target | A temperature management system for neonatal and pediatric patients, to control and monitor the patient's temperature through conductive heat transfer. For use on children up to 50 kg body mass | Cold & Heat generating device, mattresses/wraps, interconnecting hoses, accessories | Automatic, manual and full treatment modes | 12°C to 39°C | 32°C to 38°C Patient temperature accuracy within ± 0.3°C | Yes through approved calibrated probes (YSI 400 series compatible) connected to the Tecotherm Neo device | ● High and Low fluid temperature alarms ● High and Low patient temperature alarms ● Fluid flow and low reservoir alarms ● For patient temperature deviation of more than 0.5°C from set temperature | ● Fluid is TECOmed fluid (water-based) ● Fluid flow rate 0.3 liters/min to 0.5 liters/min. ● Instant feedback monitoring of infant's temperature every 2 seconds
Note: This document is for informational purposes only and is not intended to replace individual clinical decision-making, which is the sole and independent responsibility of the practitioner. HealthTrust expressly disclaims any liability for treatment decisions. Please direct any questions or comments to clinical.research@healthtrustpg.com, or to your supplier representative. ©2021 HealthTrust. All Rights Reserved.
|  |  |  | Initial or Update | Date | Completed By | Changes Made
|  |  |  | Updated | 03.2021; 04.2021 | FP | Formatting
|  |  |  | Initial | 02.2021 | HAM | Created